## Slide 1: Supplement 246 – DICOMweb Modality Procedure Step Services WG27 Final Text
September 2025

## Slide 2: Work item 2023-10-C – DICOMweb Modality Services
Introduction
The DICOM Standard defines several services. Two of these are targeted towards modalities, namely the Modality Worklist service (see PS3.4, Annex K, more specifically K.6.1) and the Modality Performed Procedure Step service (see PS3.4, Annex F, more specifically F.7-F.9). Currently, these services are defined using DIMSE.
Limitations of Current Standard
Both the Modality Worklist service and the Modality Performed Procedure Step service are not yet available in DICOMweb. This limits a) the uptake of DICOMweb for modalities and b) the support of workflow services for modalities that are (intended to be) part of a web-based ecosystem.
Description of Proposal
Add the Modality Worklist and the Modality Performed Procedure Step services to DICOMweb, in principle based on the existing DICOMweb Worklist service (UPS-RS; see PS3.18, section 11). This would boil down to creating an informative annex and any normative changes needed if gaps are discovered.

## Slide 3: Overview of Services
Supplement 246 – DICOMweb Modality Procedure Step Services contains two new services:
Modality Scheduled Procedure Step Service, mimicking DIMSE’s Basic Worklist Management Service (PS3.4, Annex K).
Modality Performed Procedure Step Service, mimicking DIMSE’s Procedure Step SOP Classes (PS3.4, Annex F), where the DIMSE notifications are not covered.
  To achieve notification-like behavior, it is recommended to mimic the approach taken in IHE’s Scheduled Workflow integration profile, where the Actor Modality Performed Procedure Step Manager forwards the creation and updating of Modality Performed Procedure Steps to other Actors that are interested in progress.
The principle of basing the Modality Performed Procedure Step Services on the existing DICOMweb Worklist Service (UPS-RS) has been relaxed to more closely mirror the DIMSE services.

## Slide 4: Modality Scheduled Procedure Step Service Resources and Transactions
Resource | URI Template | Description
Modality Scheduled Procedure Steps | /modality-scheduled-procedure-steps | The collection of Modality Scheduled Procedure Steps managed by the origin server.
Transaction Name | Method | Payload |  | Description
 |  | Request | Success Response | 
Search | GET | none | dataset according to PS3.4, Table K.6-1 | Searches for Modality Scheduled Procedure Steps

## Slide 5: Modality Scheduled Procedure Step Service DIMSE Relation and SYNTAX
Transaction | Operation | Reference | DIMSE Service
Search | Query Worklist | PS3.4, K.4 | C-FIND
Transaction | Request Syntax
Search | GET SP /modality-scheduled-procedure-steps?{&match*}{&includefield}{&fuzzymatching}{&offset}{&limit} SP version CRLF Accept: 1#media-type CRLF *(header-field CRLF) CRLF

## Slide 6: Modality Performed Procedure Step Service RESources and Transactions
Resource | URI Template | Description
Modality Performed Procedure Step | /modality-performed-procedure-steps/{mppsUID} | A Modality Performed Procedure Step.
Transaction Name | Method | Payload |  | Description
 |  | Request | Success Response | 
Create | POST | dataset according to PS3.4, Table F.7.2-1 (N-CREATE) | none | Creates a new Modality Performed Procedure Step
Update | POST | dataset according to PS3.4, Table F.7.2-1 (N-SET) | none | Updates the target Modality Performed Procedure Step
Retrieve | GET | none | dataset according to PS3.4, Table F.8.2-1 | Retrieves the target Modality Performed Procedure Step

## Slide 7: Modality Performed Procedure Step Service DIMSE Relation
Transaction | Operation | Reference | DIMSE Service
Create | Create MPPS Instance | PS3.4, F.7.2.1 | N-CREATE
Update | Set MPPS Information | PS3.4, F.7.2.2 | N-SET
Retrieve | Get MPPS Information | PS3.4, F.8.2.1 | N-GET

## Slide 8: Modality Performed Procedure Step Service SYNTAX
Transaction | Request Syntax
Create | POST SP /modality-performed-procedure-steps/{mppsUID} SP version CRLF Accept: 1#media-type CRLF *(header-field CRLF) CRLF Payload
Update | POST SP /modality-performed-procedure-steps/{mppsUID}/update SP version CRLF Accept: 1#media-type CRLF *(header-field CRLF) CRLF payload
Retrieve | GET SP /modality-performed-procedure-steps/{mppsUID}{?includefield*} SP version CRLF Accept: 1#media-type CRLF *(header-field CRLF) CRLF

## Slide 9: Modality Procedure STEP Services Support of Transactions
Service | Transaction | Support
N.A. | Retrieve Capabilities | Required
Modality Scheduled Procedure Step Service | Search | Required
Modality Performed Procedure Step Service | Create | Required
 | Update | Required
 | Retrieve | Optional

## Slide 10: Example of a Dual-Headed Server in Mixed Eco-System Basic Worklist Service / Modality Scheduled Procedure Step Service
A PACS might implement both DIMSE and DICOMweb interfaces to the same procedure step resources.
Note that although some messages are drawn next to each other, this does not mean they need to happen at the same time. It is only done to indicate semantic equivalence.

## Slide 11: Example of a Dual-Headed Server in Mixed Eco-System Modality Performed Procedure Step Service

## Slide 12: References
This presentation, the supplement, and the examples (and much more) can be found at
https://github.com/krotz-dieter/dicomweb-dmwl-mpps